## Slide 1
2008/2009 SWGDRUG ACCOMPLISHMENTS
Drug Enforcement Administration
Office of Forensic Sciences
sponsored by the
National Institute of Standards
and Technology
and the

## Slide 2
SWGDRUG
Scientific Working Group for the Analysis of Seized Drugs

## Slide 3: SWGDRUG UPDATE
The core committee voted to adopt the uncertainty document in July 2008
What Next? – Three Active Subcommittees
  1) Uncertainty Subcommittee
    Develop Supplemental Documents
  2) Education and Training Subcommittee
    Devise Comprehensive Training Program
  3) Editorial/Communications Subcommittee
    Revise/Edit Current SWGDRUG Recommendations

## Slide 4: UNCERTAINTY SUBCOMMITTEE
Working on Supplementary Documents to include “real world” examples
  Supplementary Documents
    Intended to be a resource for those implementing recommendations
    Not all inclusive, many ways to implement recommendations
Goals
  To cover as many laboratory situations as possible and make them clear and concise
  Qualitative and quantitative methods addressed
  Extensive references provided

## Slide 5
Two general approaches
  “Bottom-up” uncertainty budget
  “Top-down” incorporating method validation and continuing QA/QC such as control charts
Examples adapted from working laboratory examples
Examples are being/will be vetted through professional metrologists
Goal is to have documents ready for community review this summer
UNCERTAINTY SUBCOMMITTEE

## Slide 6
In this method, balance tolerances and known performance specifications are used
  Utilizing propagation of error approach
    Valuable for identifying most and least significant contributing sources to uncertainty
UNCERTAINTY BUDGET FOR WEIGHING EXAMPLE

## Slide 7
Determine net weight of a white powder received in a plastic bag using a top loading balance (max. capacity 4100g). The following conditions apply:
  The operator is competent on the use of the balance
  The balance is calibrated and certified
  The balance is performing within manufacturer specifications
  The balance operates at an ambient temperature varying ±5 °C
The weight recorded for the powder, determined by placing the material inside a tared weighing dish, is 30.03 grams
EXAMPLE SCENARIO

## Slide 8
Readability
Repeatability
Linearity
Sensitivity
Sample loss in transfer: The uncertainty associated with sample loss is, for practical purposes, indeterminate and irrelevant
FACTORS TO CONSIDER

## Slide 9
Factors | Value (x) | Standard uncertainty (u)* | Distribution Type* | % Contribution (Index)
Readability | 0.01 g* | = 0.00288675 g | Rectangular | 15.46%
Repeatability(s)1 | 0.01 g* | 0.01 g | Normal | 53.54%
Linearity | 0.02 g* | = 0.0057735 g | Rectangular | 30.91%
Sensitivity | 6 ppm/°C* | =1.73 x 10-5 g | Rectangular | 0.09%
1 Repeatability is listed as a standard deviation as denoted by “(s)”*manufacturer’s literature
* References are provided in the document explaining the underlying theory and provides details regarding the calculation methods
UNCERTAINTY BUDGET TABLE

## Slide 10
The combined standard uncertainty can be expressed mathematically as:
where u is the standard uncertainty and uc is combined standard uncertainty. u(sensitivity) et.al. are not included in the combined uncertainty due to their minimal % contribution index.
The combined standard uncertainty for the above is calculated as:
CALCULATION OF COMBINED STANDARD UNCERTAINTY

## Slide 11
The expanded uncertainty can be expressed mathematically as:
U = k*uc
Where U is the expanded uncertainty and k is the coverage factor.
Using a coverage factor (k) = 2.0 (confidence level of approximately 95%), U = 2*0.01190 g = 0.02380 g
Using a coverage factor (k) = 3.0 (confidence level of approximately 99.7%), U = 3*0.01190 g = 0.03571 g
CALCULATION OF EXPANED UNCERTAINTY

## Slide 12
Results
Net Weight: 30.03 grams
Confidence Range: ± 0.02380 grams
*Confidence range refers to a 95% confidence level
or
Net Weight: 30.03 grams
Confidence Range: ± 0.03571 grams
*Confidence range refers to a 99.7% confidence level
EXAMPLE SCENARIO

## Slide 13
Second example covers additive weights (multiple exhibits); similar level of detail
Quantitative methods are in early form but in essence capture uncertainty by means of method validation, controls and other QC protocols
Address single lab and multi-lab organizations
WHAT NEXT?

## Slide 14
Task – Devise comprehensive training program
Coordinate efforts with ENFSI Drugs Working Group
Develop on-line program
  Downloadable
  Hypertext Linking
  References
EDUCATION AND TRAINING SUBCOMMITTEE

## Slide 15
4.2 Topic areas in the training program will include, as a minimum, the following:
  Relevant background information on drugs of abuse (e.g., status of control and chemical and physical characteristics)
  Techniques, methodologies and instrumentation utilized in the examination of seized drug samples and related materials
  Quality assurance
  Expert /Court testimony and legal requirements
  Laboratory policy and procedures (e.g., sampling, uncertainty, evidence handling, safety and security) as they relate to the examination of seized drug samples and related materials.
EXISTING TRAINING PROGRAM

## Slide 16
1. Drugs of Abuse-General Knowledge
2. Drug Analysis
3. Forensic Context
EDUCATION AND TRAINING

## Slide 17
1. Drugs of Abuse-General Knowledge
  1.1 Classification of Drugs
  1.2 Drug Chemistry
  1.3 Street Knowledge
  1.4 Clandestine Lab Chemistry/Investigation
  1.5 Drug Profiling
EDUCATION AND TRAINING

## Slide 18
2. Drug Analysis
  2.1 General Principles of Analytical Techniques
  2.2 Category C tests
  2.3 Category B tests
  2.4 Category A tests
  2.5 Analytical Schemes
EDUCATION AND TRAINING

## Slide 19
3. Forensic Content
  3.1 Evidence handling and security
  3.2 Legal framework
  3.3 Accreditation
EDUCATION AND TRAINING

## Slide 20
Goal – Revise existing document to:
  Bring up to date
  Correct sections in conflict
  Identify any grammatical issues
  Add references to new documents
  Clarify recommendations as appropriate
EDITORIAL COMMITTEE

## Slide 21
2 Education and experience for analysts
  Removed:
    by January 1, 2005, a minimum of five (5) years practical experience in the area of seized drug analysis, and demonstrated competency following the completion of a formal, documented training program and post training competency assessment.
  Revised:
    All newly recruited analysts shall have at least a bachelor’s degree (or equivalent, generally a three to four year post-secondary or tertiary degree) in a natural science or in other sciences relevant to the analysis of seized drugs. The degree program shall include lecture and associated laboratory classes in general, organic and analytical chemistry.
EDITORIAL COMMITTEE

## Slide 22
3 Continuing professional development
  As Written:
    3.1 Twenty contact hours of training every year. Contact is defined as face-to-face interaction with an instructor or trainer in a classroom or laboratory setting. It does not include self-paced learning or distance education where the instructor has no active interaction with the student.
  As Revised:
    3.1 Twenty hours of training every year. Training can be either face-to-face interaction with an instructor, distance learning or computer based.
EDITORIAL COMMITTEE

## Slide 23
11 Analytical method validation and verification
  As Written:
    11.1 Method validation is required to demonstrate that methods are suitable for their intended purpose.
      11.1.1 For qualitative analysis, the parameters that need to be checked are selectivity, limit of detection and reproducibility.
      11.1.2 Minimum acceptability criteria should be described along with means for demonstrating compliance.
      11.1.3 Validation documentation is required.
    11.2 Laboratories adopting methods validated elsewhere should verify these methods and establish their own limits of detection and reproducibility.
  As Revised:
    Method validation is required to demonstrate that methods are suitable for their intended purpose (see PART IV B – Validation).
EDITORIAL COMMITTEE

## Slide 24
Added hyperlinks and references to UNCERTAINTY throughout document
Added hyperlinks and references to VALIDATION throughout document
Part IIIB Drug Identification, Category A to include: X-Ray Diffractometry
EDITORIAL COMMITTEE

## Slide 25
3.1 … Use second technique …
3.1.2 When sample size allows, the second technique should be applied on a separate sampling for quality assurance reasons. When sample size is limited, additional measures should be taken to assure that the results correspond to the correct sample.
3.4 In cases where hyphenated techniques are used (e.g. gas chromatography-mass spectrometry, liquid chromatography-diode array ultraviolet spectroscopy), they will be considered as separate techniques provided that the results from each are used.
PART IIIB DRUG IDENTIFICATION

## Slide 26
Problem
  If two samplings important, why have different procedure for trace samples
  Misinterpretation of 3.4, hyphenated techniques do not offer second sampling
Solution
  Revise section to emphasize quality assurance step
    Second sampling
    Procedural blank
    Witnessing
PART IIIB DRUG IDENTIFICATION

## Slide 27
CORE COMMITTEE
  DEA – Nelson Santos (Chair)
  Secretariat – Scott Oulton (non-voting)
  FBI - Eileen Waninger
  ASCLD – Garth Glassburg
  NIST - Susan Ballou
  ASTM and NEAFS- Jack Mario
  Educator – Dr. Chris Tindall
  Educator – Dr. Suzanne Bell

## Slide 28
CORE COMMITTEE
  CAC & NWAFS - Jerry Massetti
  MAFS - Richard Paulas
  MAAFS - Linda Jackson
  SAFS – Christian Matchett
  Toxicology – Dr. Robert Powers

## Slide 29
CORE COMMITTEE
  Canada - Richard Laing
  Japan – Mr. Osamu Ohtsuru
  United Kingdom - Dr. Sylvia Burns
  Australia - Catherine Quinn
  Germany - Dr. Udo Zerell
  ENFSI - Dr. Michael Bovens
  UNODC - Dr. Iphigenia Naidis

## Slide 30: THANK YOU
  Visit Us At
  www.swgdrug.org